## MFT Hardware Configurations
Scope of Offer: |  | Describe offer in terms of Item 5.1 in the specification document. AB schedule to be completed with reference to this offer.
Manufacturer: |  | XYZ
Model: |  | MF3PH
Table 1 - Summary of Hardware specification required - 240-102634558 Clause 5.1
Item | General Hardware capability |  | QTY | Notes
a) | Number of voltage channels? |  | 4
b) | Number of Neutrals? |  | 2 | Usually a neutral is grouped with a 3 phase set, e.g. 4 channels and two neutrals will indicated L1-L2-L3-N and L1-N (6 plug outputs)
c) | Can Voltage channels convert to current channels? |  | No
d) | Is the rating of converted Voltage channel equivalent to the standard current channels? |  | NA | Not Applicable
e) | If No, please define in exception document. |  | NA | Not Applicable
f) | Number of current channels? |  | 3 | L1-L2-L3-N
g) | Number of Neutrals? |  | 1
h) | Number of binary inputs |  | 10
I) | Number of binary outputs |  | 4
J) | Number of DC inputs (Transducer Testing) |  | 2
k) | Auxiliary DC power supply |  | 1
l) | Low level outputs |  | No
k) | Transducer and Tariff meter test capable? |  | No | As per specification
Table 2 - Summary of Hardware specification required - 240-102634558 Clause 5.2
Item | General Hardware capability |  | QTY | Notes
a) | Number of voltage channels?
b) | Number of Neutrals?
c) | Can Voltage channels convert to current channels? |  | No
d) | Is the rating of converted Voltage channel equivalent to the standard current channels? |  | NA | Not Applicable
e) | If No, please define in exception document. |  | NA | Not Applicable
f) | Number of current channels? |  | 3 | L1-L2-L3-N
g) | Number of Neutrals? |  | 1
h) | Number of binary inputs |  | 10
I) | Number of binary outputs |  | 4
J) | Number of DC inputs (Transducer Testing) |  | 2
k) | Auxiliary DC power supply |  | 1
l) | Low level outputs |  | No
k) | Transducer and Tariff meter test capable? |  | No | As per specification
Table 3 - Summary of Hardware specification required - 240-102634558 Clause 5.3
Item | General Hardware capability |  | QTY | Notes
a) | Number of voltage channels?
b) | Number of Neutrals?
c) | Can Voltage channels convert to current channels? |  | No
d) | Is the rating of converted Voltage channel equivalent to the standard current channels? |  | NA | Not Applicable
e) | If No, please define in exception document. |  | NA | Not Applicable
f) | Number of current channels? |  | 3 | L1-L2-L3-N
g) | Number of Neutrals? |  | 1
h) | Number of binary inputs |  | 10
I) | Number of binary outputs |  | 4
J) | Number of DC inputs (Transducer Testing) |  | 2
k) | Auxiliary DC power supply |  | 1
l) | Low level outputs |  | No
k) | Transducer and Tariff meter test capable? |  | No | As per specification
Table 4 - Summary of Hardware specification required - 240-102634558 Clause 5.4
Item | General Hardware capability |  | QTY | Notes
a) | Number of voltage channels?
b) | Number of Neutrals?
c) | Can Voltage channels convert to current channels? |  | No
d) | Is the rating of converted Voltage channel equivalent to the standard current channels? |  | NA | Not Applicable
e) | If No, please define in exception document. |  | NA | Not Applicable
f) | Number of current channels? |  | 3 | L1-L2-L3-N
g) | Number of Neutrals? |  | 1
h) | Number of binary inputs |  | 10
I) | Number of binary outputs |  | 4
J) | Number of DC inputs (Transducer Testing) |  | 2
k) | Auxiliary DC power supply |  | 1
l) | Low level outputs |  | No
k) | Transducer and Tariff meter test capable? |  | No | As per specification

## M - Mandatory Criteria
METERING MANDATORY CRITERIA - 240-170000773
Mandatory Criteria | Enquiry Returnable | Comply (Yes/No) | Comments
Are completed Technical Schedules A&B, including any deviations, submitted? | Completed and signed A&B Technical Schedules as in the excel spreadsheet.
Are completed A&B Support Requirements Schedules, including any deviations, submitted? | Completed and signed A&B Technical Schedules as in the excel spreadsheet.
Are Deviations to the Referenced Technical Standard submitted with the associated declaration signed by the duly authorised representative? | Completed and signed Deviation Schedules as in the excel spreadsheet.
Are completed Annexure A and Annexure B of 240-170000797 submitted? | Completed and signed Annexures A and B of 240-170000797.
Are all information supplied in English? | Documents, brochures, supporting documents supplied.
Threshold. Should the Supplier fail to meet ANY ONE of the above requirements they will be disqualified.
Company Name: |  | Date
Representative Name |  | Signature

## A1-Technical Requirements AB
TECHNICAL REQUIREMENTS A&B SCHEDULE FOR MULTI-FUNCTION CONTROL PLANT TEST SET SPECIFICATION 240-170000773 METERING TECHNICAL REQUIREMENTS
Schedule A: Purchaser specific requirements
Schedule B: Guarantees and technical particulars of equipment offered
When completing the Schedule of Requirements A/B Suppliers are requested to take cognisance of the following: • It is expected of Suppliers to state clearly, for each clause that requires a statement of compliance in the Schedule of Requirements A/B, either “Comply” or “Do not comply”. • If a clause in the Schedule of Requirements A/B requires a statement of compliance and additional information e.g. a description etc., Suppliers are requested to state clearly either “Comply – followed by any information deemed appropriate by the Supplier” or “Do not comply – followed by any information deemed appropriate by the Supplier”. • If a clause in the Schedule of Requirements A/B requires information only, e.g. describe application etc., Suppliers are at liberty to provide any information deemed appropriate. Supporting information, such as specification sheets, instruction manuals, test reports and certificates should be referenced where applicable. Electronic copies of the information referenced shall be provided and can be hyperlinked where possible to assist with the evaluation. The Supplier shall provide detailed evidence for each answer specifically stating the exact location of the evidence in the supporting documentation i.e. document name and reference, page number, paragraph number It is critical that references to supporting documentation (document title, page number and clause / paragraph no.) that confirm compliance shall be clearly indicated on the Technical Schedules in the "Reference" column.
Item | Clause of 240-170000773 | Description | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
1 | 2.2.1 | Normative references referred to by Supplier | Yes/No |  | NA
2 | 2.2.2 | Informative reference referred to by Supplier | Yes/No |  | NA | 1
3 | 3.1 | Multifunction test set casing and carry handle | Conform, list exceptions |  |  | 3
4 | 3.2 | Front panel connections. | Conform, list exceptions |  |  | 3
5 | 3.2.1 | Voltage outputs | Conform, list exceptions |  |  | 3
6 | 3.2.1 | Detail seperately with exception list the protection provided if any, should a output inadvertently be energised | Yes, detailed in exception list |  |  | 2
7 | 3.2.2 | Current outputs | Conform, list exceptions |  |  | 3
8 | 3.2.2 | Detail seperately with exception list the protection provided if any should a output inadvertently be energised | Yes, detailed in exception list |  |  | 2
9 | 3.2.4 | Low level current and voltage outputs | Yes/No |  |  | 1
10 | 3.2.7 | Analogue DC inputs | Conform, list exceptions |  |  | 3
11 | 3.2.8 | Auxiliary power supply | Conform, list exceptions |  |  | 3
12 | 3.2.9 | Input and output terminals | Conform, list exceptions |  |  | 3
13 | 3.3 | Rear panel | Conform, list exceptions |  |  | 2
14 | 3.4 | Power supply -SANS 164-0 -SANS 603201 | Conform, list exceptions |  |  | 3
15 | 3.5 | Ethernet ports -Test set link to PC for software control -Time synchronisation via PTP -Testing of IEC61850 Publish and subscribe to GOOSE -Testing IEC61850 Sampled Values -Copper ethernet -Fibre optic ethernet -Power over ethernet (PoE) | Conform, list exceptions |  |  | 3
16 | 3.5.1 | Other communication ports -Optional software control via PC interface | Conform, list exceptions |  |  | 2
17 | 3.6 | Indication to be given as to the deemed test set vulnerability to cyber-attacks, and any preventative measures which are in place. Preference is for the test set to be secured from cyber-attacks to minimise the MFT as the entry point for any malicious attacks. | Conform, list exceptions |  |  | 3
18 | 4.1 | Power supply | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
19 | a) | Nominal Voltage | 230 VAC |  |  | 3
20 | b) | Permissible Voltage Range | 85 V - 265 VAC |  |  | 3
21 | c) | Nominal Frequency | 50 Hz / 60 Hz |  |  | 3
22 | d) | Permissible Frequency Range | 45 Hz - 65 Hz |  |  | 3
23 | e) | Power Consumption | Max 3450VA at 230VAC |  |  | 3
24 | f) | Rated Current | 15A at 230VAC |  |  | 3
25 | g) | Connection | Standard AC sockets: Connector (SANS 60320-1, C13) and Inlet (SANS 60320-1, C14) Other SANS approved connectors and inlets may be considered. |  |  | 3
26 | 4.2 | Dimensions and weight | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
27 | a) | Volume | max. 45 liter |  |  | 2
28 | b) | Weight | max. 20 kg |  |  | 2
29 | 4.3 | Environment | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
30 | a) | Operating Temperature | 0 °C - 50 °C |  |  | 3
31 | b) | Storage Temperature | -25 °C - 70 °C |  |  | 3
32 | c) | Humidity | 5 % - 95 % RH, non-condensing |  |  | 3
33 | 4.4 | Certificates | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
34 | a) | Vibration | IEC60068-2-6 |  |  | 3
35 | b) | Shock | IEC60068-2-27 |  |  | 3
36 | c) | CE Mark | Yes or equivalent. |  |  | 3
37 | d) | EMC | The product adheres to the electromagnetic compatibility (EMC) Directive 2004 / 108 / EC (CE conform) |  |  | 3
38 | e) | EMC - Emissions | IEC 61326 Class A; IEC61000-6-4; SANS 61000-3-2/3; FCC Subpart B of Part 15 Class A |  |  | 3
39 | f) | EMC - Immunity | IEC 61326; IEC61000-6-2; IEC61000-4-2/3/4/5/6/11 |  |  | 3
40 | g) | Safety | The product adheres to the low voltage Directive 2006 / 95 / EC (CE conform). IEC61010-1: Safety requirements for electrical equipment for measurement, control, and laboratory use; Insulation of PC and SELV Interfaces complies with SANS 60950-1 (Also EN) |  |  | 3
41 | h) | Independent Test Laboratory | Copies of test certificates from independent test laboratories proving compliance with the above standards are to be supplied together with the tender |  |  | 3
42 | 4.5 | Hardware | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
43 | 4.5.1 | General | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
44 | a) | All functions should be combined in one hardware unit | Conform, list exceptions |  |  | 3
45 | b) | The unit shall be of a robust and sturdy construction | Conform, list exceptions |  |  | 3
46 | c) | PC card design: Wire jumpers on a single printed circuit board are not permissible | Conform, list exceptions |  |  | 3
47 | d) | Electronic components: No potentiometers are allowed | Conform, list exceptions |  |  | 3
48 | e) | No moving elements or elements that are susceptible to damage, i.e. controls elements or displays on the face plate are permissible | Conform, list exceptions |  |  | 3
49 | 4.5.2 | Output Amplifiers | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
50 | a) | The amplifier stages are to be fully electronic, i.e., not via transformer | Conform, list exceptions |  |  | 3
51 | b) | All current amplifiers to be fully protected against any open-circuit-, overload-, overburden- and over-temperature- condition. Any such condition is to be immediately displayed in all active software modules or test set specific controller. Except for an over-temperature condition an automatic shutdown of the amplifiers is not permissible. The withstand capability of the outputs when inadvertently connected to an external supply to be defined. | Conform, list exceptions |  |  | 3
52 | c) | All voltage amplifiers to be fully protected and immune against any short-circuit-, overload-, overburden- and over-temperature- condition. Any such condition is to be immediately displayed in all active software modules or test set specific controller. Except for an over-temperature condition an automatic shutdown of the amplifiers is not permissible | Conform, list exceptions |  |  | 3
53 | d) | All amplifiers to use linear amplification elements and to be dc-coupled. The neutrals of the voltage and current amplifiers shall be galvanically isolated (up to 2kV) from the ground/chassis of the test set. This is for the safety of the test set operating personnel and will ensure that the test set casing does not inadvertently subject the operator to any electrical shock. | Conform, list exceptions |  |  | 3
54 | e) | Ability to generate DC and AC signals | Conform, list exceptions |  |  | 3
55 | f) | The amplifiers, low-level outputs, the measurement inputs and the main power supply to be galvanically isolated from each other and earth (2kV insulation voltage) | Conform, list exceptions |  |  | 3
56 | 4.5.3 | Voltage Amplifiers | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
57 | a) | Setting range | 4x 0 - 300Vrms (L-N), 3x 0 - 520Vrms (L-L) |  |  | 3
58 | b) | Single phase operation | 1x 0 - 600Vrms (L-L) |  |  | 3
59 | c) | Output power | 3x 100VA at 100 - 300V (L-N) or |  |  | 3
60 | d) | Output power | 1x 200VA at 100 - 300V (L-N) or |  |  | 3
61 | e) | Output power | 1x 275VA at 200 - 600V (L-L) |  |  | 3
62 | f) | Minimum load current | 3x 1Arms; 1x 2Arms |  |  | 3
63 | g) | Accuracy | error < 0.05% of reading (0 - 300V) + 0.02% of range |  |  | 3
64 | h) | THD+N | < 0.015% |  |  | 3
65 | I) | Phase error | <0.02° |  |  | 3
66 | j) | Connection | 4mm Banana sockets |  |  | 3
67 | 4.5.4 | Current Amplifiers | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
68 | a) | Setting range | 6x 0 - 30Arms |  |  | 3
69 | b) | Setting range | 3x 0 - 60Arms |  |  | 3
70 | c) | Single phase operation | 1x 0 - 125Arms |  |  | 3
71 | d) | Output power | 6x 430VA at 25A |  |  | 3
72 | e) | Output power | 3x 860VA at 50A |  |  | 3
73 | f) | Output power | 1x 1000VA at 80A (LL-LN) or |  |  | 3
74 | g) | Output power | 1x1740VA at 25A (L-L-L-L) |  |  | 3
75 | h) | Maximum compliance voltage | 6x 35Vpk; (higher is better.) |  |  | 3
76 | I) | Maximum compliance voltage | 1x 140Vpk (higher is better.) |  |  | 3
77 | j) | Accuracy | error < 0.05% of reading (0 - 32A) + 0.02% or range |  |  | 3
78 | k) | THD+N | < 0.015% |  |  | 3
79 | l) | Phase error | <0.05° |  |  | 3
80 | m) | Independent amplifiers | The six current amplifiers are to be independent from the four voltage amplifiers. The neutrals between the two current output groups are to be galvanically isolated (up to 2kV) to allow for series connection of the two groups in order to boost the compliance voltage. |  |  | 3
81 | n) | Connection | 4mm Banana sockets |  |  | 3
82 | 4.5.9 | Analogue low level measuring inputs for transducer testing | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
83 | a) | Direct current range | Range 1: 0 - ±1 mA |  |  | 2
84 | b) |  | Range 2: 0 - ±20 mA |  |  | 3
85 | c) | Direct voltage range | 0 - ±10 V |  |  | 3
86 | d) | Max. error | < 0.003% |  |  | 3
87 | e) | Connection | 4mm Banana sockets |  |  | 3
88 | 4.5.13 | Interface to PC | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
89 | a) | Interface to the IBM compatible PC via Ethernet interface or USB | Detail |  |  | 3
90 | b) | 10/100Mbit/s copper (auto sensing, auto crossover) via RJ45 connector alternatively USB. (10/1000Mbit/S shall be preferred and is essential where the test set processes sampled values) | Detail |  |  | 3
91 | c) | Two Ethernet communications ports to support communication on a process bus (IEC61850-8-1 (GOOSE), IEC61850-9-2 (Sampled Values) and UCA2.0 AND station bus at the same time. PTP to IEEE 1588 and IEC61850-9-3. | No |  |  | 3
92 | d) | Communications card to support IEEE 1588 / Precision Time Protocol (PTP) and IEC61850-9-3 to synchronize the test set to a PTP enabled Grandmaster clock in the substation for End-to-end tests / Synchrophasor tests. | Detail |  |  | 3
93 | e) | Centronics parallel port (LPT) and/or serial RS232 ports are not permissible as modern laptops do not provide such ports. | Detail |  |  | 2
94 | 4.6 | Software | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
95 | 4.6.1 | General | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
96 | a) | Windows 10 (32bit & 64bit) / 7(32bit & 64bit) software, upgradeable to the latest version available from Microsoft. Long filenames, tool tip help, context sensitive menu function (right mouse click) and an integrated help browser shall be provided. All future versions of Windows operating systems within the lifespan of the hardware to be supported. | Yes/No |  |  | 3
97 | b) | All software functions, options and actions should be easily available by click of a button and/or shortcut key to avoid having to navigate through complicated menu structures and having to drill through multiple menu levels. The Microsoft ribbon-based menu structure used in Office 365 is an example of sorting all functions, options and actions and making all easily accessible in the right place. | Yes/No |  |  | 3
98 | c) | No programming to be necessary to test an application - entry of setting parameters to be all that is required to set up and perform a test. | Yes/No |  |  | 3
99 | d) | Future expansions in functionality by means of software updates. Firmware updating to be handled by the software, i.e. exchange of any hardware components is not permissible. | Yes/No |  |  | 3
100 | e) | Generation of reports on paper or file. All graphics and text to be printable. | Yes/No |  |  | 3
101 | f) | Test report shall be configurable to include custom information in graphical format (e.g. logos, wiring diagrams) and text format (tester, date of test, substation, reminders). Once a test report is generated, it shall be time stamped and will not be editable. | Yes/No |  |  | 3
102 | g) | The test set software should be able to import relay settings records from OEM relay setting software currently used by the Purchaser, in a format compatible with the test set. | Yes/No |  |  | 3
103 | h) | It shall be possible to test multi-function relays / scheme with one test routine / document. | Yes/No |  |  | 3
104 | i) | Test sequence shall allow a pause in the test sequence at pre-defined points, by popping up a custom instruction dialogue (with or without audible warning), instructing the user to change either a setting on the relay, change of hard wiring or to record a specific measurement / status from the relay / scheme. | Yes/No |  |  | 3
105 | j) | All testing to be in closed loop. | Yes/No |  |  | 3
106 | k) | Online Pass/ Fail assessment for ALL tests. This is particularly important for automatic testing. | Yes/No |  |  | 3
107 | l) | Full automatic testing shall be possible, i.e. without launching various test modules manually. | Yes/No |  |  | 3
108 | m) | Allow for tolerances to be adjusted in the test templates, for example an adjustable range from 0% to 15%. | Yes/No |  |  | 3
109 | n) | MFT Software shall not be hindered by the use of anti-virus software eg. Mcafee/Norton. | Yes/No |  |  | 3
110 | 4.6.15 | Power Quality Signal Generator to NRS048-2 / SANS/IEC61000-4-30; SANS/IEC61000-4-7; BS EN/IEC 61000-4-15 and IEC62586 | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
111 | a) | Testing of Power Quality measurement devices as per the NRS 048 standard. | Yes/No |  |  | 3
112 | 4.6.16 | On-line multi-meter measurement function | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
113 | a) | Up to 10 independent inputs | Yes/No |  |  | 2
114 | b) | Software to display amplitudes and phase angles of ac voltages and currents, symmetrical components of voltages and currents, line to line voltages, frequencies, power (active, reactive and apparent) and cos(phi) independently for each input. | Yes/No |  |  | 2
115 | c) | All ac measurements to be true RMS. | Yes/No |  |  | 2
116 | d) | On-line vector diagram for voltages and currents as well as power. | Yes/No |  |  | 2
117 | e) | On-line measurement of Vdc, Idc and dc power for each input. | Yes/No |  |  | 2
118 | f) | On-line display of harmonics measured both numerically (in a table per harmonic frequency) and graphically in a bar graph. | Yes/No |  |  | 2
119 | 4.7 | Energy Meter Testing | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
120 | a) | The Multi-function Test Set for energy meter testing can be a separate test set or option/module within the Secondary plant MFT |  |  |  | 3
121 | b) | The MFT shall be able to handle the testing of meters for both three phase three wire and three phase four wire configurations for current transformers operated meters both at 1A and 5A |  |  |  | 3
122 | c) | The MFT for energy meters shall be made up of the following components: • Source • Reference/working standard • Scanning head • Test software |  |  |  | 1
123 | d) | The MFT shall be able to test energy meters using the existing network load or at specific load points utilising the MFT source. |  |  |  | 3
124 | e) | The MFT shall facilitate both manual and automatic testing of energy meters. |  |  |  | 3
125 | 4.7.1 | Source | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
126 |  | THE MFT source for energy meter testing shall comply with the requirements specified in the following sections of the standard:
127 | a) | 4.1 Power Supply |  |  |  | 1
128 | b) | 4.2 Dimensions and Weight |  |  |  | 1
129 | c) | 4.3 Environment |  |  |  | 1
130 | d) | 4.4 Certificates |  |  |  | 1
131 | e) | 4.5.1 General Hardware |  |  |  | 1
132 | f) | 4.5.2 Output Amplifiers |  |  |  | 1
133 | g) | 4.5.3 Voltage Amplifiers |  |  |  | 1
134 | h) | 4.5.4 Current amplifiers |  |  |  | 1
135 | i) | 4.5.9 Analogue low level measuring inputs for transducer testing |  |  |  | 1
136 | j) | 4.5.13 interface to PC |  |  |  | 1
137 | 4.7.2 | Reference/Working Standard | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
138 | a) | The reference/working standard shall be capable of performing single and three-phase measurements with and accuracy class of 0.05% for active energy and 0.2% for reactive energy. The reference/working standard shall have the required accuracy to test class 0.1 S metering devices. The Purchaser and SANS/IEC standards that the Supplier shall prove the suitability of the reference/working standard to test a class 0.1 S energy meter are as follows: |  |  |  | 3
139 | b) | 240-77224537: Standard for Calibration and Limits of Errors for Single and Three Phase Energy Meters |  |  |  | 1
140 | c) | SANS/IEC 62052-11: Electricity metering equipment (a.c) – General requirements, tests and test conditions Part 11: Metering equipment |  |  |  | 1
141 | d) | IEC 62053-22: Electricity metering equipment (a.c) – Particular requirements Part 22: Static meters for active energy (classes 0,1S, 0.2S and 0,5S) |  |  |  | 1
142 | e) | SANS/IEC 62053-24: Electricity metering equipment (a.c) – Particular requirements Past 24: Static meters for reactive energy at fundamental frequency (classes 0,5S, 1S, 1, 2 and 3) |  |  |  | 1
143 | f) | SANS 474: Electrical metering – Standard requirements |  |  |  | 1
144 | g) | It is preferred that the reference/working standard be a separate module or detachable from the MFT to facilitate its calibration. |  |  |  | 1
145 | h) | Voltage measuring range | 0V to 300V (L-N) |  |  | 3
146 | i) | Current measuring range | 0 to 10A |  |  | 3
147 | j) | Frequency range | 45 Hz to 65 Hz |  |  | 3
148 | k) | Measurement mode for single phase meters | 1 phase 2 wire active / reactive |  |  | 1
149 | l) | Measurement mode for three phase meters | 1 phase 2 wire active / reactive |  |  | 3
150 | m) | Measurement mode for three phase meters | 3 phase 4 wire active / reactive |  |  | 3
151 | n) | Measurement mode for three phase meters | 3 phase 3 wire active / reactive |  |  | 3
152 | 4.7.2.1 | Reference/Working Standard Power Supply (If separate unit) | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
153 | a) | The power supply for the reference/working standard, if the reference/working standard is a separate unit shall comply with the requirements specified in section 4.1 of the standard |  |  |  | 1
154 | 4.7.2.2 | Reference/Working Standard Functionality | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
155 |  | Reference working standard shall provide for the following functions
156 | a) | Measurement of electrical parameters including voltage, current, active, reactive and apparent power, active, reactive and apparent energy, frequency, phase angle, power factor and harmonics |  |  |  | 3
157 | b) | Error calculation utilising the pulses counted from the meter and provide an indication of the percentage error of the meter under test. |  |  |  | 3
158 | c) | Vector diagram, harmonic analysis and wave form display |  |  |  | 3
159 | d) | Current instrument transformer tests: ratio and burden |  |  |  | 3
160 | e) | Voltage instrument transformer tests: ratio, burden and voltage drop |  |  |  | 3
161 | 4.7.3 | Test Software | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
162 |  | The test software shall comply with the general requirements specified in section 4.6.1 of the standard with the following additions:
163 | a) | The software shall allow generating and storing test programs. It shall be possible to adapt the programs to different meter types, nominal values and ranges of current and voltage. |  |  |  | 3
164 | b) | The software shall allow automatic or manual (step-by-step) execution of test programs. |  |  |  | 3
165 | c) | It shall control, display and log the test parameters, such as connection mode, voltages, current, frequency, phase angles, power factor, power and energy. |  |  |  | 3
166 | d) | The software shall monitor the operation of the MFT during execution of the program and if parameters are outside the acceptable limits, give a warning or alarm and/or abort the program. |  |  |  | 3
167 | e) | The software shall provide information on the status, for example “Initialisation”, “Measurement running”, “Waiting for input from operator”, “Aborted”, “Finished”, etc. |  |  |  | 3
168 | f) | The software shall automatically calculate the errors and evaluate the results of the tests for each test point. It shall display and print the results and store them with proper identifiers for further processing. |  |  |  | 3
169 | g) | The software shall be able to generate vector diagrams, harmonic analysis and wave form displays, print and store them in graphical format with proper identifiers. |  |  |  | 3
170 | h) | The software shall be able to perform instrument transformer tests (ratio, burden and voltage drop test for voltage instrument transformers), print and store the results with proper identifiers. |  |  |  | 3
171 | 4.7.4 | Scanning Head | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
172 | a) | The receiver diode of the scanning head shall be capable of sensing optical output according to SANS/IEC 62052 -11 clauses 5.11.1 and 5.11.2 |  |  |  | 3
173 | 4.7.5 | Error Calculation System | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
174 |  | The error calculation system shall be able to count pulses and to calculate and indicate the percentage error of the energy meter under test. The error calculation system receives pulses from the scanning head and compares those with the pulses received from the reference/working standard |  |  |  | 3
175 | a) | The system shall provide an automatic assessment of the results (pass/fail) and indicate the error along with the sign (+ or -) |  |  |  | 3
176 | b) | A reset function shall be available allowing to reset the error indication in the case the error measurement is incorrect, for any reason |  |  |  | 2
177 | c) | The system shall provide the parameters of the error calculation process for verification purposes |  |  |  | 3
178 | d) | Pulse frequency range, which the system shall be able to count | 0 to 1 kHz |  |  | 3
179 | e) | Resolution of error indication | Minimum 0,01% |  |  | 3
180 | f) | Indication of error | +/- indication, numbers along with % |  |  | 3
181 | g) | Accuracy | +/- 1 digit of the lowest significant digit |  |  | 3
182 | 4.8 | Software Licensing | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
183 | a) | A software license shall cover all operating system versions where the functionality is similar. The future of desktop operating systems is hanging in the balance; end users should have experience with other operating systems such as mobile operating systems. Investment in software should be for the function not the operating system type. A software license shall not depend on a hardware dongle to be used. The License shall be imbedded in the software and the Purchaser’s license shall be embedded in the software package such that any employee of the Purchaser may control a test set of another employee using his / her own computer and test set software, it is accepted modules are registered per test set and that certain functions may not be available to another user. Dedicated software for use with say an IEC61850 package shall also not use a dongle. The license shall be applied such that a user can transfer the license to other computers. | Yes/No |  |  | 3
184 | 4.9 | Local Support | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
185 | a) | Hotline support hours | Local telephonic support to be offered in the hours of 07h00 - 20h00 South African time. |  |  | 3
186 | b) | Electronic communication support | Any form of electronic communication to be offered with a guaranteed turnaround time of 1 business day, e.g. MS Teams |  |  | 3
187 | c) | Number of units in operation in South Africa | 10 units in SA (separate list of users with contact details to be provided) |  |  | 3
188 | d) | No of back-up units available in South Africa | 2 (Full Specification Test Set) |  |  | 3
189 | e) | Calibration facility available in South Africa | Yes |  |  | 3
190 | f) | Guaranteed turnaround time for repairs and calibration | < 14 working days |  |  | 3
191 | g) | Product training offered in South Africa | a) Standard training to be offered as a registered course with Eskom Academy of Learning. b) In-house / On-site training to be offered on request |  |  | 3
192 | h) | Warranty on Hardware | Five years from date of delivery |  |  | 3
193 | 5 | Test Set configurations | State which test set is offered? (Y/N) - |  | Suppliers reference/Evidence | Weight
194 | 5.5 | Energy Meter Test Set | Y/N |  |  | 3
195 | 5.5.1 | Universal three phase test set with refence/working standard | Y/N |  |  | 2
196 | 5.5.2 | Standalone universal three phase test set | Y/N |  |  | 1
197 | 5.5.3 | Standalone reference/working standard | Y/N |  |  | 1
198 | 6 | Software | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
199 |  | Supplier Complies to all requirements listed under this clause? | Yes/No and provide details. |  |  | 3
200 | 6.1 | Basic Software | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
201 | d) | Energy meter testing as defined in section 4.7. | Yes/No |  |  | 3
202 | Note | (Additional functionality may be listed it) | Detail in exception report |  |  | 3
203 | 6.2 | Intermediate software - includes Basic software items and the additional listed functionality. | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
204 | e) | Transducer (single and three phase) testing as defined in item 4 | Yes/No |  |  | 3
205 | i) | Energy meter testing as defined in section 4.7 | Yes/No |  |  | 3
206 | 6.3 | Advanced Advanced modules - This software package shall be separate from the intermediate package, and procured as required by the end user. | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
207 | a) | All software functionality as indicated in the basic and intermediated. | Yes/No |  |  | 3
208 | e) | On Line multi-meter function as defined in item 4 | Yes/No |  |  | 2
209 | f) | Power Quality measurement device testing based on NRS048-2 / SANS/IEC61000-4-30; SANS/IEC61000-4-7; BS EN/IEC 61000-4-15 and IEC62586. In essence NRS 048 is the standard followed in South Africa and the Supplier shall state whether the test module is fully compliant with the NRS 048 standard. | Yes/No |  |  | 3
210 | 7 | Accessories | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
211 | 7.1 | Standard accessory package with a MFT: | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
212 | a) | Each test set to be equipped with a set of 2m; 2 mm² (minimum length) test leads sufficient to populate all inputs/outputs | Yes/No |  |  | 3
213 | b) | A combination test lead with the 4 current and voltage leads, thus 8 leads in total clearly marked for current and voltage application and suitably colour coded. This lead shall be long enough to plug into the current and voltage channels and reach the current and voltage test points on a typical protection or metering scheme. | Yes/No |  |  | 3
214 | c) | Connectors for each lead to allow a sound connection to the devices under test. | Yes/No |  |  | 3
215 | d) | A quality carry bag made of a durable material with carrying strap or other quality case. The bag or container shall easily house all the leads and accessories. | Yes/No |  |  | 3
216 | e) | A hardened transport case with wheels to facilitate easy transport and sufficiently sturdy to be used for multiple shipping events. Lockable. | Yes/No |  |  | 2
217 | f) | The required communication cables | Yes/No |  |  | 3
218 | g) | Manual, the manual may be part of the software module or a software copy. The manual shall be fully detailed. The manual is to cover both the hardware and the software. The manuals may be two separate software documents. The Test set manuals shall preferably be imbedded in the test set control software and have an auto update facility when connected to the Internet. | Yes/No |  |  | 3
219 | h) | Should a software module require a specific accessory to complete or execute a test the test set package shall come equipped with the required accessory and be accordingly priced. | Yes/No |  |  | 3
220 | i) | A calibration certificate. See the details under the calibration section. | Yes/No |  |  | 3
221 | 7.2 | Examples of other Accessories or Services | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
222 | a) | The listed examples are currently used by the Purchaser, however ALL useful software modules and accessories on offer to be listed and priced in terms of the contract pricing structure and period. | Yes/No |  |  | 3
223 | b) | Clamps, crocodile clips, terminal connectors. | Yes/No |  |  | 3
224 | c) | Clip on current transducers for test sets with current measurement/recording capabilities. | Yes/No |  |  | 3
225 | f) | Sensors to accommodate measurement of all types of meters. (Pulsing LED signals) | Yes/No |  |  | 3
226 | g) | Additional communications cables and other specialised cables. | Yes/No |  |  | 3
227 | h) | Software modules | Yes/No |  |  | 3
228 | i) | Calibration service to SANAS requirement, priced for all test set offerings including return shipping. | Yes/No |  |  | 3
229 | j) | Repair / Calibration return shipping, however this is usually included in the repair / calibration cost and shall be the preferred method. | Yes/No |  |  | 3
230 | l) | Switch boxes | Yes/No |  |  | 3
231 | m) | Test set controllers with on board software (Not a PC) | Yes/No |  |  | 3
232 | p) | Hardware upgrades | Yes/No |  |  | 3
233 | q) | Any other useful accessory or software module | Yes/No |  |  | 3
Company Name: |  |  |  | Date
Representative Name |  |  |  | Signature

## A2-Support Requirements AB
A&B SCHEDULE FOR MULTI-FUNCTION CONTROL PLANT TEST SET SPECIFICATION 240-170000773 METERING SUPPORT REQUIREMENTS
Schedule A: Purchaser specific requirements
Schedule B: Guarantees and technical particulars of equipment offered
When completing the Schedule of Requirements A/B Suppliers are requested to take cognisance of the following: • It is expected of Suppliers to state clearly, for each clause that requires a statement of compliance in the Schedule of Requirements A/B, either “Comply” or “Do not comply”. • If a clause in the Schedule of Requirements A/B requires a statement of compliance and additional information e.g. a description etc., Suppliers are requested to state clearly either “Comply – followed by any information deemed appropriate by the Supplier” or “Do not comply – followed by any information deemed appropriate by the Supplier”. • If a clause in the Schedule of Requirements A/B requires information only, e.g. describe application etc., Suppliers are at liberty to provide any information deemed appropriate. Supporting information, such as specification sheets, instruction manuals, test reports and certificates should be referenced where applicable. Electronic copies of the information referenced shall be provided and can be hyperlinked where possible to assist with the evaluation. The Supplier shall provide detailed evidence for each answer specifically stating the exact location of the evidence in the supporting documentation i.e. document name and reference, page number, paragraph number It is critical that references to supporting documentation (document title, page number and clause / paragraph no.) that confirm compliance shall be clearly indicated on the Technical Schedules in the "Reference" column.
Clause of 240-170000773 | Description | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
8.1 | Calibration
a) | Ability of Supplier to confirm fully to Clause 8.1 | Comply - Provide evidence of a calibration certificate conforming to clause 8,1 |  |  | 3
8.2 | Repair | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Ability of Supplier to confirm fully to Clause 8.2 | Comply |  |  | 3
8.3 | Annual Report | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Ability of Supplier to confirm fully to Clause 8.3 | Comply - Provide evidence of a database available and capability to modify metadata. |  |  | 2
9 | Training | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Ability of Supplier to confirm fully to Clause 9 | Comply |  |  | 3
9.1 | Basic Training | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Is the relevant training material available and ability to conform to requirements? | Y/N and is Provide Training Structure/Brochure from Company. |  |  | 3
9.2 | Intermediate Training | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Is the relevant training material available and ability to conform to requirements? | Y/N and is Provide Training Structure/Brochure from Company. |  |  | 3
9.3 | Metering and Measurements Training | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Is the relevant training material available and ability to conform to requirements? | Y/N and is Provide Training Structure/Brochure from Company. |  |  | 3
9.4 | Advanced Training | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Is the relevant training material available and ability to conform to requirements? | Y/N and is Provide Training Structure/Brochure from Company. |  |  | 3
9.5 | Specialist Training | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Is the relevant training material available and ability to conform to requirements? | Y/N and is Provide Training Structure/Brochure from Company. |  |  | 3
9.6 | Trainee Evaluation | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Ability of supplier to conform | Y/N |  |  | 3
10 | Support | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Local repair and calibration with a turnaround time typically better than 14 days. | State Compliance: (Comply/Do Not Comply) Provide Evidence. |  |  | 3
b) | Formalised training courses as described. Training shall take place at or near the Purchaser‘s offices around the country, when sufficient student numbers justify this; or virtual training as required. | State Compliance: (Comply/Do Not Comply) Provide Evidence. |  |  | 3
c) | Technical support shall be available via email and telephonically or any other form of electronic communication by technical competent support staff who is within 2 hours of South African time zone. Technical competent would mean personnel with tertiary Electrical engineering qualifications and having hands on experience testing protection IED’s that are used in the Purchaser’s division (Transmission.) | State Compliance: (Comply/Do Not Comply) Provide Evidence. |  |  | 3
e) | Lastly the Supplier shall provide proof of the required support functions. This will be a list of technical staff and a short CV of each member, courses provided over the past years and to which companies. Physical examples of automated templates and the end customer. | State Compliance: (Comply/Do Not Comply) Provide Evidence. |  |  | 3
11 | Automated Test Templates | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Capability and capacity to develop automated testing templates on request. | State Compliance: (Comply/Do Not Comply) Provide Evidence. |  |  | 3
b) | Compatibility with the Purchaser’s existing automated test templates is required. Alternatively, Suppliers can submit proposals to migrate existing automated test templates to a format that is compatible with their proposed products. | State Compliance:or submit proposals |  |  | 3
|  |  | Total Risk and Support Score |  | 47
Company Name: |  |  |  | Date
Representative Name |  |  |  | Signature

## B1 - Demonstration
PRACTICAL DEMONSTRATION FOR METERING MULTI-FUNCTION CONTROL PLANT TEST SET SPECIFICATION 240-170000773
Number | Demonstration | Weighting | Supplier Score | Comments
1 | Accuracy testing of energy meters | 20 | 2
2 | Accuracy testing of transducers | 20 | 2
3 | Curent instrument transformer testing - ratio and burden | 10 | 2
4 | Voltage instrument transformer testing - ratio, burden and voltage drop | 10 | 2
5 | Ability of test set to provide test reports usable to Eskom, eg in pdf | 20 | 2
6 | Ability of test set to generate vector diagrams, harmonic analysis and waveform display | 20 | 2
Company Name: |  |  | Date
Eskom Evaluator Names |  |  | Eskom Evaluator Signature

## Deviation Schedule
Deviation Schedule
Any deviations from the 240-170000773 specification shall be listed below with reasons for deviation. In addition, evidence shall be provided that the proposed deviation will at least be more cost effective than that specified by Eskom.
Item | Clause | Proposed Deviation
Company Name
Representative Name
Date
Signature