## MFT Hardware Configurations
Scope of Offer: |  | Describe offer in terms of Item 5.1 in the specification document. AB schedule to be completed with reference to this offer.
Manufacturer: |  | XYZ
Model: |  | MF3PH
Table 1 - Summary of Hardware specification required - 240-102634558 Clause 5.1
Item | General Hardware capability |  | QTY | Notes
a) | Number of voltage channels? |  | 4
b) | Number of Neutrals? |  | 2 | Usually a neutral is grouped with a 3 phase set, e.g. 4 channels and two neutrals will indicated L1-L2-L3-N and L1-N (6 plug outputs)
c) | Can Voltage channels convert to current channels? |  | No
d) | Is the rating of converted Voltage channel equivalent to the standard current channels? |  | NA | Not Applicable
e) | If No, please define in exception document. |  | NA | Not Applicable
f) | Number of current channels? |  | 3 | L1-L2-L3-N
g) | Number of Neutrals? |  | 1
h) | Number of binary inputs |  | 10
I) | Number of binary outputs |  | 4
J) | Number of DC inputs (Transducer Testing) |  | 2
k) | Auxiliary DC power supply |  | 1
l) | Low level outputs |  | No
k) | Transducer and Tariff meter test capable? |  | No | As per specification
Table 2 - Summary of Hardware specification required - 240-102634558 Clause 5.2
Item | General Hardware capability |  | QTY | Notes
a) | Number of voltage channels?
b) | Number of Neutrals?
c) | Can Voltage channels convert to current channels? |  | No
d) | Is the rating of converted Voltage channel equivalent to the standard current channels? |  | NA | Not Applicable
e) | If No, please define in exception document. |  | NA | Not Applicable
f) | Number of current channels? |  | 3 | L1-L2-L3-N
g) | Number of Neutrals? |  | 1
h) | Number of binary inputs |  | 10
I) | Number of binary outputs |  | 4
J) | Number of DC inputs (Transducer Testing) |  | 2
k) | Auxiliary DC power supply |  | 1
l) | Low level outputs |  | No
k) | Transducer and Tariff meter test capable? |  | No | As per specification
Table 3 - Summary of Hardware specification required - 240-102634558 Clause 5.3
Item | General Hardware capability |  | QTY | Notes
a) | Number of voltage channels?
b) | Number of Neutrals?
c) | Can Voltage channels convert to current channels? |  | No
d) | Is the rating of converted Voltage channel equivalent to the standard current channels? |  | NA | Not Applicable
e) | If No, please define in exception document. |  | NA | Not Applicable
f) | Number of current channels? |  | 3 | L1-L2-L3-N
g) | Number of Neutrals? |  | 1
h) | Number of binary inputs |  | 10
I) | Number of binary outputs |  | 4
J) | Number of DC inputs (Transducer Testing) |  | 2
k) | Auxiliary DC power supply |  | 1
l) | Low level outputs |  | No
k) | Transducer and Tariff meter test capable? |  | No | As per specification
Table 4 - Summary of Hardware specification required - 240-102634558 Clause 5.4
Item | General Hardware capability |  | QTY | Notes
a) | Number of voltage channels?
b) | Number of Neutrals?
c) | Can Voltage channels convert to current channels? |  | No
d) | Is the rating of converted Voltage channel equivalent to the standard current channels? |  | NA | Not Applicable
e) | If No, please define in exception document. |  | NA | Not Applicable
f) | Number of current channels? |  | 3 | L1-L2-L3-N
g) | Number of Neutrals? |  | 1
h) | Number of binary inputs |  | 10
I) | Number of binary outputs |  | 4
J) | Number of DC inputs (Transducer Testing) |  | 2
k) | Auxiliary DC power supply |  | 1
l) | Low level outputs |  | No
k) | Transducer and Tariff meter test capable? |  | No | As per specification

## M - Mandatory Criteria
PROTECTION MANDATORY CRITERIA - 240-170000773
Mandatory Criteria | Enquiry Returnable | Comply (Yes/No) | Comments
Are completed Technical Schedules A&B, including any deviations, submitted? | Completed and signed A&B Technical Schedules as in the excel spreadsheet.
Are completed A&B Support Requirements Schedules, including any deviations, submitted? | Completed and signed A&B Technical Schedules as in the excel spreadsheet.
Are Deviations to the Referenced Technical Standard submitted with the associated declaration signed by the duly authorised representative? | Completed and signed Deviation Schedules as in the excel spreadsheet.
Are completed Annexure A and Annexure B of 240-170000797 submitted? | Completed and signed Annexures A and B of 240-170000797.
Are all information supplied in English? | Documents, brochures, supporting documents supplied.
Threshold. Should the Supplier fail to meet ANY ONE of the above requirements they will be disqualified.
Company Name: |  | Date
Representative Name |  | Signature

## A1- A&B Technical Schedules
A&B SCHEDULE FOR MULTI-FUNCTION CONTROL PLANT TEST SET SPECIFICATION 240-170000773 PROTECTION TECHNICAL REQUIREMENTS
Schedule A: Purchaser specific requirements
Schedule B: Guarantees and technical particulars of equipment offered
When completing the Schedule of Requirements A/B Suppliers are requested to take cognisance of the following: • It is expected of Suppliers to state clearly, for each clause that requires a statement of compliance in the Schedule of Requirements A/B, either “Comply” or “Do not comply”. • If a clause in the Schedule of Requirements A/B requires a statement of compliance and additional information e.g. a description, reasons etc., Suppliers are requested to state clearly either “Comply – followed by any information deemed appropriate by the Supplier” or “Do not comply – followed by any information deemed appropriate by the Supplier”. • If a clause in the Schedule of Requirements A/B requires information only, e.g. describe application etc., Suppliers are at liberty to provide any information deemed appropriate. Supporting information, such as specification sheets, instruction manuals, test reports and certificates should be referenced where applicable. Electronic copies of the information referenced shall be provided and can be hyperlinked where possible to assist with the evaluation. The Supplier shall provide detailed evidence for each answer specifically stating the exact location of the evidence in the supporting documentation i.e. document name and reference, page number, paragraph number It is critical that references to supporting documentation (document title, page number and clause / paragraph no.) that confirm compliance shall be clearly indicated on the Technical Schedules in the "Reference" column.
Item | Clause of 240-170000773 | Description | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
1 | 2.2.1 | Normative references referred to by Supplier | N/A
2 | 2.2.2 | Informative reference referred to by Supplier | N/A
3 | 3.1 | Multifunction test set casing and carry handle | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
4 | 3.2 | Front panel connections. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
5 | 3.2.1 | Voltage outputs | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
6 | 3.2.1 | Detail seperately with exception list the protection provided if any, should a output inadvertently be energised | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
7 | 3.2.2 | Current outputs | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
8 | 3.2.2 | Detail seperately with exception list the protection provided if any should a output inadvertently be energised | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
9 | 3.2.3 | Synchronising of Test set current and voltage channels across multiple test set units | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
10 | 3.2.4 | Low level current and voltage outputs | Comply provide evidence, where Do not Comply provide reasons. |  |  | 1
11 | 3.2.5 | Binary inputs | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
12 | 3.2.6 | Binary outputs | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
13 | 3.2.7 | Analogue DC inputs | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
14 | 3.2.8 | Auxiliary power supply | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
15 | 3.2.9 | Input and output terminals | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
16 | 3.3 | Rear panel | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
17 | 3.4 | Power supply -SANS 164-0 -SANS 603201 | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
18 | 3.5 | Ethernet ports -Test set link to PC for software control -Time synchronisation via PTP -Testing of IEC61850 Publish and subscribe to GOOSE -Testing IEC61850 Sampled Values -Copper ethernet -Fibre optic ethernet -Power over ethernet (PoE) | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
19 | 3.5.1 | Other communication ports -Optional software control via PC interface | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
20 | 3.6 | Indication to be given as to the deemed test set vulnerability to cyber-attacks, and any preventative measures which are in place. Preference is for the test set to be secured from cyber-attacks to minimise the MFT as the entry point for any malicious attacks. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
21 | 4.1 | Power supply | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
22 | a) | Nominal Voltage | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
23 | b) | Permissible Voltage Range | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
24 | c) | Nominal Frequency | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
25 | d) | Permissible Frequency Range | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
26 | e) | Power Consumption | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
27 | f) | Rated Current | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
28 | g) | Connection | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
29 | 4.2 | Dimensions and weight | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
30 | a) | Volume | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
31 | b) | Weight | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
32 | 4.3 | Environment | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
33 | a) | Operating Temperature | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
34 | b) | Storage Temperature | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
35 | c) | Humidity | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
36 | 4.4 | Certificates | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
37 | a) | Vibration | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
38 | b) | Shock | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
39 | c) | CE Mark | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
40 | d) | EMC | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
41 | e) | EMC - Emissions | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
42 | f) | EMC - Immunity | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
43 | g) | Safety | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
44 | h) | Independent Test Laboratory | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
45 | 4.5 | Hardware | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
46 | 4.5.1 | General | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
47 | a) | All functions should be combined in one hardware unit | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
48 | b) | The unit shall be of a robust and sturdy construction | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
49 | c) | PC card design: Wire jumpers on a single printed circuit board are not permissible | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
50 | d) | Electronic components: No potentiometers are allowed | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
51 | e) | No moving elements or elements that are susceptible to damage, i.e. controls elements or displays on the face plate are permissible | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
52 | 4.5.2 | Output Amplifiers | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
53 | a) | The amplifier stages are to be fully electronic, i.e., not via transformer | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
54 | b) | All current amplifiers to be fully protected against any open-circuit-, overload-, overburden- and over-temperature- condition. Any such condition is to be immediately displayed in all active software modules or test set specific controller. Except for an over-temperature condition an automatic shutdown of the amplifiers is not permissible. The withstand capability of the outputs when inadvertently connected to an external supply to be defined. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
55 | c) | All voltage amplifiers to be fully protected and immune against any short-circuit-, overload-, overburden- and over-temperature- condition. Any such condition is to be immediately displayed in all active software modules or test set specific controller. Except for an over-temperature condition an automatic shutdown of the amplifiers is not permissible | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
56 | d) | All amplifiers to use linear amplification elements and to be dc-coupled. The neutrals of the voltage and current amplifiers shall be galvanically isolated (up to 2kV) from the ground/chassis of the test set. This is for the safety of the test set operating personnel and will ensure that the test set casing does not inadvertently subject the operator to any electrical shock. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
57 | e) | Ability to generate DC and AC signals | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
58 | f) | The amplifiers, low-level outputs, the measurement inputs and the main power supply to be galvanically isolated from each other and earth (2kV insulation voltage) | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
59 | 4.5.3 | Voltage Amplifiers | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
60 | a) | Setting range | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
61 | b) | Single phase operation | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
62 | c) | Output power | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
63 | d) | Output power | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
64 | e) | Output power | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
65 | f) | Minimum load current | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
66 | g) | Accuracy | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
67 | h) | THD+N | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
68 | I) | Phase error | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
69 | j) | Connection | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
70 | 4.5.4 | Current Amplifiers | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
71 | a) | Setting range | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
72 | b) | Setting range | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
73 | c) | Single phase operation | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
74 | d) | Output power | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
75 | e) | Output power | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
76 | f) | Output power | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
77 | g) | Output power | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
78 | h) | Maximum compliance voltage | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
79 | I) | Maximum compliance voltage | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
80 | j) | Accuracy | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
81 | k) | THD+N | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
82 | l) | Phase error | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
83 | m) | Independent amplifiers | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
84 | n) | Connection | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
85 | 4.5.5 | Low-level signal generators | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
86 | a) | 6 additional analogue low level signals to control external amplifiers or to test relays with low-level input, e.g. from Rogowski coils, linear voltage or linear current sensors must be provided | Comply provide evidence, where Do not Comply provide reasons. |  |  | 1
87 | b) | Full simulation of Rogowski coil signal (i.e. first order differential of signal) to be performed in hardware. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 1
88 | c) | Together with 10 internally used signal generators the system should provide 16 independent signal channels in total | Comply provide evidence, where Do not Comply provide reasons. |  |  | 1
89 | d) | Output range | Comply provide evidence, where Do not Comply provide reasons. |  |  | 1
90 | e) | Accuracy | Comply provide evidence, where Do not Comply provide reasons. |  |  | 1
91 | f) | Overload Protection | Comply provide evidence, where Do not Comply provide reasons. |  |  | 1
92 | g) | The quantities displayed in the software must be scalable for primary or secondary voltages (or currents). | Comply provide evidence, where Do not Comply provide reasons. |  |  | 1
93 | h) | Couple to other auxiliary devices or possible synchronising function. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 1
94 | 4.5.6 | Signal Generation | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
95 | a) | All outputs to be continuously and independently adjustable in amplitude, phase (0 to +/- 360°) and frequency. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
96 | b) | Able to generate continuous sine waves with a frequency between 10 and 1kHz and to generate transient files with a bandwidth from dc up to 3 kHz. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
97 | c) | Frequency error to be less than 0.5 ppm. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
98 | d) | Phase error to be less than 0.02° | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
99 | 4.5.7 | Binary inputs | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
100 | a) | Number of inputs | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
101 | b) | Mode/Trigger criteria | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
102 | c) | Max. Input Voltage | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
103 | d) | Max. error | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
104 | e) | Counting function | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
105 | f) | Connection | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
106 | 4.5.8 | Analogue inputs for measurement | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
107 | a) | Number | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
108 | b) | Input ranges | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
109 | c) | Accuracy | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
110 | d) | Bandwidth | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
111 | e) | Sampling frequency | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
112 | f) | Overload protection | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
113 | g) | Recording buffer | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
114 | h) | Connection | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
115 | 4.5.9 | Analogue low level measuring inputs for transducer testing | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
116 | a) | Direct current range | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
117 | b) |  | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
118 | c) | Direct voltage range | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
119 | d) | Max. error | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
120 | e) | Connection | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
121 | 4.5.10 | Auxiliary DC supply to power test objects | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
122 | a) | Range | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
123 | b) |  | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
124 | c) |  | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
125 | d) | Power | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
126 | e) | Max. error | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
127 | 4.5.11 | Binary output contacts | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
128 | a) | Number | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
129 | b) | Breaking capacity | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
130 | c) | Type | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
131 | 4.5.12 | Time synchronisation | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
132 | a) | The test set should be able to synchronise to a GPS clock via a 1pps or IRIG B signal. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
133 | b) | Timing accuracy for 1pps / IRIG-B: 1 μs. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
134 | c) | Maximum distance between GPS receiver and test equipment for 1pps signal: 40m. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
135 | d) | Test set needs to be synchronised to the IEEE 1588 Precision Time Protocol (PTP) via Ethernet and conform to IEC61850-9-3, IEEE C37.238-2011 and IEC61588. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
136 | e) | Timing accuracy for PTP: 100ns. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
137 | f) | Maximum distance between GPS receiver and test equipment for PTP: 2000m. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
138 | g) | IEEE 1588 / PTP clock to be powered over Ethernet (PoE), i.e. no separate power supply should be necessary. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
139 | 4.5.13 | Interface to PC | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
140 | a) | Interface to the IBM compatible PC via Ethernet interface or USB | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
141 | b) | 10/100Mbit/s copper (auto sensing, auto crossover) via RJ45 connector alternatively USB. (10/1000Mbit/S shall be preferred and is essential where the test set processes sampled values) | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
142 | c) | Two Ethernet communications ports to support communication on a process bus (IEC61850-8-1 (GOOSE), IEC61850-9-2 (Sampled Values) and UCA2.0 AND station bus at the same time. PTP to IEEE 1588 and IEC61850-9-3. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
143 | d) | Communications card to support IEEE 1588 / Precision Time Protocol (PTP) and IEC61850-9-3 to synchronize the test set to a PTP enabled Grandmaster clock in the substation for End-to-end tests / Synchrophasor tests. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
144 | e) | Centronics parallel port (LPT) and/or serial RS232 ports are not permissible as modern laptops do not provide such ports. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
145 | 4.6 | Software | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
146 | 4.6.1 | General | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
147 | a) | Windows 10 (32bit & 64bit) / 7(32bit & 64bit) software, upgradeable to the latest version available from Microsoft. Long filenames, tool tip help, context sensitive menu function (right mouse click) and an integrated help browser shall be provided. All future versions of Windows operating systems within the lifespan of the hardware to be supported. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
148 | b) | All software functions, options and actions should be easily available by click of a button and/or shortcut key to avoid having to navigate through complicated menu structures and having to drill through multiple menu levels. The Microsoft ribbon-based menu structure used in Office 365 is an example of sorting all functions, options and actions and making all easily accessible in the right place. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
149 | c) | No programming to be necessary to test an application - entry of setting parameters to be all that is required to set up and perform a test. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
150 | d) | Future expansions in functionality by means of software updates. Firmware updating to be handled by the software, i.e. exchange of any hardware components is not permissible. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
151 | e) | Generation of reports on paper or file. All graphics and text to be printable. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
152 | f) | Test report shall be configurable to include custom information in graphical format (e.g. logos, wiring diagrams) and text format (tester, date of test, substation, reminders). Once a test report is generated, it shall be time stamped and will not be editable. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
153 | g) | The test set software should be able to import relay settings records from OEM relay setting software currently used by the Purchaser, in a format compatible with the test set. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
154 | h) | It shall be possible to test multi-function relays / scheme with one test routine / document. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
155 | i) | Test sequence shall allow a pause in the test sequence at pre-defined points, by popping up a custom instruction dialogue (with or without audible warning), instructing the user to change either a setting on the relay, change of hard wiring or to record a specific measurement / status from the relay / scheme. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
156 | j) | All testing to be in closed loop. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
157 | k) | Online Pass/ Fail assessment for ALL tests. This is particularly important for automatic testing. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
158 | l) | Full automatic testing shall be possible, i.e. without launching various test modules manually. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
159 | m) | Allow for tolerances to be adjusted in the test templates, for example an adjustable range from 0% to 15%. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
160 | n) | MFT Software shall not be hindered by the use of anti-virus software eg. Mcafee/Norton. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
161 | 4.6.2 | Manual control | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
162 | a) | Direct entry of actual relay settings into test software. Test specific parameters (e.g. set I1 to 110% of I>>) shall be automatically re-adjusted according to the entered relay settings. | Comply, where Do not Comply provide reasons. |  |  | 3
163 | b) | Manual and independent adjustment of amplitude, phase angle and frequency for all generator outputs. | Comply, where Do not Comply provide reasons. |  |  | 3
164 | c) | Graphical display of natural voltages and currents in a vector diagram. | Comply, where Do not Comply provide reasons. |  |  | 3
165 | d) | Direct entry sequence components and graphical display in a vector diagram | Comply, where Do not Comply provide reasons. |  |  | 3
166 | e) | Direct entry of impedances and graphical display in a R/X diagram | Comply, where Do not Comply provide reasons. |  |  | 3
167 | f) | Direct entry of power and graphical display in P/Q diagram | Comply, where Do not Comply provide reasons. |  |  | 3
168 | g) | Ramping and stepping of any of the above quantities: all phases at the same time. | Comply, where Do not Comply provide reasons. |  |  | 3
169 | h) | Synchronized switching of multiple variables at the same instance in time. | Comply, where Do not Comply provide reasons. |  |  | 3
170 | i) | Timing functions for pick-up and drop-out measurements. | Comply, where Do not Comply provide reasons. |  |  | 3
171 | j) | On-line Reporting Function | Comply, where Do not Comply provide reasons. |  |  | 3
172 | k) | Synchronise generator outputs to any third-party signal, e.g. mains frequency. | Comply, where Do not Comply provide reasons. |  |  | 3
173 | l) | Full control of voltage and current. Full control of amplitude and frequency. Any test required by a MFT shall be possible using manual control. | Comply, where Do not Comply provide reasons. |  |  | 3
174 | 4.6.3 | State sequencer function | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
175 | a) | Direct entry of actual relay settings into test software. Test specific parameters (e.g. set I1 to 110% of I>>) shall be automatically re-adjusted according to the entered relay settings. | Comply, where Do not Comply provide reasons. |  |  | 3
176 | b) | Manual testing. | Comply, where Do not Comply provide reasons. |  |  | 3
177 | c) | Fully automatic testing. | Comply, where Do not Comply provide reasons. |  |  | 3
178 | d) | Ability to generate test sequences from any number of states. Each state consists of any combination of voltage, current, frequency and any binary output state. | Comply, where Do not Comply provide reasons. |  |  | 3
179 | e) | Graphical display of natural voltages and currents in a vector diagram. | Comply, where Do not Comply provide reasons. |  |  | 3
180 | f) | Graphical display of voltages, currents and binary signals over time. | Comply, where Do not Comply provide reasons. |  |  | 3
181 | g) | Direct entry sequence components and graphical display in a vector diagram | Comply, where Do not Comply provide reasons. |  |  | 3
182 | h) | Direct entry of impedances and graphical display in a R/X diagram | Comply, where Do not Comply provide reasons. |  |  | 3
183 | i) | Direct entry of power and graphical display in P/Q diagram | Comply, where Do not Comply provide reasons. |  |  | 3
184 | j) | Define trigger conditions for each state either in fixed time or dependent upon a logical combination of the binary inputs for accurate timing measurements. | Comply, where Do not Comply provide reasons. |  |  | 3
185 | k) | Synchronization to GPS / PTP and other digital timing pulses. | Comply, where Do not Comply provide reasons. |  |  | 3
186 | l) | Automatic assessment of test results (Pass / Fail) with actual results and deviation from nominal. | Comply, where Do not Comply provide reasons. |  |  | 3
187 | 4.6.4 | Linear ramping and pulse ramping functions | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
188 | a) | Direct entry of actual relay settings into test software. Test specific parameters (e.g. set I1 to 110% of I>>) shall be automatically re-adjusted according to the entered relay settings. | Comply, where Do not Comply provide reasons. |  |  | 3
189 | b) | Manual testing. | Comply, where Do not Comply provide reasons. |  |  | 3
190 | c) | Fully automatic testing. | Comply, where Do not Comply provide reasons. |  |  | 3
191 | d) | Ability to linearly ramp up to two independent variables (e.g. voltage and frequency) at the same time, while keeping the other quantities at a defined constant value. | Comply, where Do not Comply provide reasons. |  |  | 3
192 | e) | Ability to ramp a single variable in a stepped fashion i.e. ramped quantity shall be set back to a predefined value between every step to allow the relay to reset. | Comply, where Do not Comply provide reasons. |  |  | 3
193 | f) | Graphical display of natural voltages and currents in a vector diagram. | Comply, where Do not Comply provide reasons. |  |  | 3
194 | g) | Graphical display of voltages, currents and binary signals over time. | Comply, where Do not Comply provide reasons. |  |  | 3
195 | h) | Define trigger conditions for pick-up / drop-out measurements upon a logical combination of the binary inputs. | Comply, where Do not Comply provide reasons. |  |  | 3
196 | i) | Synchronization to GPS / PTP and other digital timing pulses. | Comply, where Do not Comply provide reasons. |  |  | 3
197 | j) | Automatic assessment of test results (Pass / Fail) with actual results and deviation from nominal. | Comply, where Do not Comply provide reasons. |  |  | 3
198 | 4.6.5 | Transient playback function | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
199 | a) | IEEE COMTRADE (C37.111-1991) compatible (ASCII and binary formats). | Comply, where Do not Comply provide reasons. |  |  | 3
200 | b) | Synchronization to GPS / PTP and other digital timing pulses. | Comply, where Do not Comply provide reasons. |  |  | 3
201 | c) | Recording playback sampling rate of up to 4.8kHz (minimum). Down sampling to be supported. | Comply, where Do not Comply provide reasons. |  |  | 3
202 | d) | Graphical display of voltage and current traces as well as digital signals and relay responses. | Comply, where Do not Comply provide reasons. |  |  | 3
203 | e) | Editing of signals: Extending (repeating) and deleting portions of the recorded signal. | Comply, where Do not Comply provide reasons. |  |  | 3
204 | f) | Ability to generate composite harmonic wave shapes. | Comply, where Do not Comply provide reasons. |  |  | 3
205 | g) | Ability to edit own digital signals in the COMTRADE file. The test set software must allow existing binaries to be played back to outputs. | Comply, where Do not Comply provide reasons. |  |  | 3
206 | h) | Minimum length of Comtrade files to play back shall be 10 minutes. | Comply, where Do not Comply provide reasons. |  |  | 3
207 | 4.6.6 | Overcurrent relay testing | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
208 | a) | Direct entry of actual relay settings into test software. Test specific parameters (e.g. test points) shall be automatically re-adjusted according to the entered relay settings. | Comply, where Do not Comply provide reasons. |  |  | 3
209 | b) | Manual testing. | Comply, where Do not Comply provide reasons. |  |  | 3
210 | c) | Fully automatic testing. | Comply, where Do not Comply provide reasons. |  |  | 3
211 | d) | Testing of the operating characteristic (trip time vs. current test) for all types of fault (earth faults, phase faults, negative sequence and zero sequence faults) | Comply, where Do not Comply provide reasons. |  |  | 3
212 | e) | Display of overcurrent characteristic and test points in I/t diagram. | Comply, where Do not Comply provide reasons. |  |  | 3
213 | f) | Determine the pick-up and drop-out current for all fault types. | Comply, where Do not Comply provide reasons. |  |  | 3
214 | g) | Determine the pick-up and drop-out of directional characteristic for directional overcurrent relays. | Comply, where Do not Comply provide reasons. |  |  | 3
215 | h) | Breaker simulation: Simulate the 52a and 52b auxiliary contacts of a breaker with the binary outputs, switch the currents off at zero crossing of current after a trip signal has been received. | Comply, where Do not Comply provide reasons. |  |  | 3
216 | i) | Characteristic formulae: IEC255-4, BS142 and IEEE PC37.112-1995, I2t characteristics to be supported. Definition of custom characteristics shall be possible. | Comply, where Do not Comply provide reasons. |  |  | 3
217 | j) | Automatic assessment of test results (Pass / Fail) with actual results and deviation from nominal. | Comply, where Do not Comply provide reasons. |  |  | 3
218 | 4.6.7 | Frequency and voltage relay testing | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
219 | a) | Direct entry of actual relay settings into test software. Test specific parameters (e.g. test points) shall be automatically re-adjusted according to the entered relay settings. | Comply, where Do not Comply provide reasons. |  |  | 3
220 | b) | Manual testing. | Comply, where Do not Comply provide reasons. |  |  | 3
221 | c) | Fully automatic testing. | Comply, where Do not Comply provide reasons. |  |  | 3
222 | d) | Generation of ramps for amplitudes, phase angles and frequency. | Comply, where Do not Comply provide reasons. |  |  | 3
223 | e) | Pick-up, timing and stability tests. | Comply, where Do not Comply provide reasons. |  |  | 3
224 | f) | Graphic and tabular display of relay pick-up and drop-out vs. time. | Comply, where Do not Comply provide reasons. |  |  | 3
225 | g) | Automatic assessment of test results (Pass / Fail) with actual results and deviation from nominal. | Comply, where Do not Comply provide reasons. |  |  | 3
226 | 4.6.8 | Distance relay testing | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
227 |  | (Testing of line impedance relays and field failure characteristics on generator protection relays.) | Comply, where Do not Comply provide reasons. |  |  | 3
228 | a) | Importing / manual entry of the actual relay settings of a distance relay into the test software. A few examples of these relays are ABB REL 670/RED670/511/531; SEL 321/421; Alstom Quadromho, Optimho, Micromho; Siemens 7SA511/3, 7SL27, R3Z27; GEC YTG, PYTS; BBC LZ32. The tripping characteristic shall be generated automatically for each of the above relays. All new IED’s should be fully supported by the MFT Supplier. | Comply, where Do not Comply provide reasons. |  |  | 3
229 | b) | Impedance characteristics supported: Quadrilateral, Mho, Tomato or Lens characteristic. | Comply, where Do not Comply provide reasons. |  |  | 3
230 | c) | Manual testing. | Comply, where Do not Comply provide reasons. |  |  | 3
231 | d) | Fully automatic testing. | Comply, where Do not Comply provide reasons. |  |  | 3
232 | e) | Test the trip time at specific fault impedances. Fault impedances shall be specifiable in relation to zone reaches, e.g. 90% of Z1. Automatic assessment of whether the tested trip time has passed or failed. | Comply, where Do not Comply provide reasons. |  |  | 3
233 | f) | Check the reach of a relay by placing a fault shot at the inner and outer tolerances border (i.e. theoretical reach minus and plus the defined reach tolerance of the relay). Automatic assessment of whether the tested reach has Passed or Failed. | Comply, where Do not Comply provide reasons. |  |  | 3
234 | g) | Automatically test the characteristic in the impedance plane (R/X diagram) and/or in the time grading diagram (Z/t diagram). Automatic assessment of whether the tested reach has passed or failed. | Comply, where Do not Comply provide reasons. |  |  | 3
235 | h) | Test models supported: constant test current, constant test voltage and constant source impedance. | Comply, where Do not Comply provide reasons. |  |  | 3
236 | i) | DC offset simulation: Control of angle of fault incidence, time constant of dc offset to be calculated on-line from system parameters (R/L). | Comply, where Do not Comply provide reasons. |  |  | 3
237 | j) | The separate arc resistance algorithm is to be supported for both earth faults and phase faults (as implemented on numerical distance relays). | Comply, where Do not Comply provide reasons. |  |  | 3
238 | k) | Apply Pre-fault voltage, i.e. duration settable. | Comply, where Do not Comply provide reasons. |  |  | 3
239 | l) | Breaker simulation: Simulate the 52a and 52b auxiliary contacts of a breaker with the binary outputs; switch the currents off at zero crossing of current after a trip signal has been received. | Comply, where Do not Comply provide reasons. |  |  | 3
240 | m) | Graphical display of analogue voltages and currents plus relay contact responses vs. time. | Comply, where Do not Comply provide reasons. |  |  | 3
241 | n) | Display of injected voltages and currents on a vector diagram in natural and/or in symmetrical components. | Comply, where Do not Comply provide reasons. |  |  | 3
242 | o) | Testing of auxiliary functions: Manual close, power swing, Auto-reclose function, VT fuse fail. | Comply, where Do not Comply provide reasons. |  |  | 3
243 | p) | Automatic assessment of test results (Pass / Fail) with actual results and deviation from nominal. | Comply, where Do not Comply provide reasons. |  |  | 3
244 | 4.6.9 | Differential relay testing | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
245 |  | (Testing of transformer differential relays, line differential relays, motor / generator differential and busbar differential relays.) | Comply, where Do not Comply provide reasons. |  |  | 3
246 | a) | Importing / Manual entry of the actual relay settings of differential relay directly into the test software. A few examples of these relays are ABB RET670/RED670/SPAD 346, Reyrolle Duobias, SEL 387/487/587/787, Reyrolle Duobias M and Siemens 7UT51. The operating characteristic shall be generated automatically for each of the above relays. Any new differential algorithm that is introduced shall be accommodated in a reasonable time period; typically this should be at no cost to the purchaser. | Comply, where Do not Comply provide reasons. |  |  | 3
247 | b) | Manual testing. | Comply, where Do not Comply provide reasons. |  |  | 3
248 | c) | Fully automatic testing. | Comply, where Do not Comply provide reasons. |  |  | 3
249 | d) | Simulation of two and three winding transformers for all possible vector groups (e.g. YY0, YD1, YD11, etc.). | Comply, where Do not Comply provide reasons. |  |  | 3
250 | e) | Testing the operating characteristic (Idiff vs. Ibias) for all types of faults: earth fault, phase faults and three phase faults. | Comply, where Do not Comply provide reasons. |  |  | 3
251 | f) | The various types of Ibias formulae (Ibias = (¦Ip¦+¦Is¦)/k; Ibias =max(Ip,Is); etc; numerical zero sequence elimination and both reference side have to be supported. | Comply, where Do not Comply provide reasons. |  |  | 3
252 | g) | Testing the harmonic restraint characteristic (Idiff vs. %I harmonic) for second harmonic (inrush restraint) and fifth harmonic (over fluxing restraint). | Comply, where Do not Comply provide reasons. |  |  | 3
253 | h) | Testing of the inrush restraint characteristic for relays which utilize the gap detection technique. | Comply, where Do not Comply provide reasons. |  |  | 3
254 | i) | Testing the trip time characteristic (trip time vs. Idiff) for all types of fault. | Comply, where Do not Comply provide reasons. |  |  | 3
255 | j) | Test the stability of the relay to confirm the correctness of the vector group correction, CT mismatch correction and zero sequence elimination. | Comply, where Do not Comply provide reasons. |  |  | 3
256 | k) | Apply pre-fault current, i.e. through fault current condition. Test current and duration settable. | Comply, where Do not Comply provide reasons. |  |  | 3
257 | l) | Apply voltage in addition to six currents - according to HV or LV voltage | Comply, where Do not Comply provide reasons. |  |  | 3
258 | m) | GPS / PTP synchronized End-to-end tests to test the operating characteristic of line differential relays, i.e. by simulating either the local or remote end of a line. | Comply, where Do not Comply provide reasons. |  |  | 3
259 | n) | Automatic assessment of test results (Pass / Fail) with actual results and deviation from nominal. | Comply, where Do not Comply provide reasons. |  |  | 3
260 | 4.6.10 | Synchronizing Devices | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
261 | a) | Importing / Manual entry of actual relay settings into test software. Test specific parameters (e.g. test points) shall be automatically re-adjusted according to the entered relay settings. | Comply, where Do not Comply provide reasons. |  |  | 3
262 | b) | Manual testing. | Comply, where Do not Comply provide reasons. |  |  | 3
263 | c) | Fully automatic testing. | Comply, where Do not Comply provide reasons. |  |  | 3
264 | d) | Adjustment control mode to be tested closed loop. | Comply, where Do not Comply provide reasons. |  |  | 1
265 | e) | Graphical display of quantities in a DV/Df diagram as well as relative phase angles in a synchronoscope. | Comply, where Do not Comply provide reasons. |  |  | 3
266 | f) | Feedback signals: closing pulse and adjustment pulses (V+, V-, f+, f-). Display of adjustment controls vs. time. | Comply, where Do not Comply provide reasons. |  |  | 3
267 | g) | Automatic assessment of test results (Pass / Fail) with actual results and deviation from nominal. | Comply, where Do not Comply provide reasons. |  |  | 3
268 | 4.6.11 | Testing with GOOSE messages | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
269 | a) | All the above protective relay test modules shall support testing for IEC 61850-8-2, i.e. be able to trigger on GOOSE messages as well as simulate GOOSE messages (if need be). Test set to support subscription to Logical Nodes as per IEC 61850-7 (All sub-parts). | Comply, where Do not Comply provide reasons. |  |  | 3
270 | b) | Import of SCD, ICD of GOOSE sniffer files to configure the triggering / simulation of GOOSE messages. | Comply, where Do not Comply provide reasons. |  |  | 3
271 | c) | Simulation and subscription of up to 360 simultaneous GOOSE messages shall be possible. | Comply, where Do not Comply provide reasons. |  |  | 3
272 | d) | GOOSE performance to IEC 61850-5 and processing time to Network and to Test set better than 1mS. | Comply, where Do not Comply provide reasons. |  |  | 3
273 | e) | VLAN support (selectable priority and VLAN-ID | Comply, where Do not Comply provide reasons. |  |  | 3
274 | 4.6.13 | Testing with sampled values | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
275 | a) | All the above protective relay test modules shall support testing for UCA Guideline IEC 61850-9-2LE, IEC 61869-9 and IEC 61850-9-2, i.e. be able to simulate all voltages and currents as Sampled Value signals. | Comply, where Do not Comply provide reasons. |  |  | 3
276 | b) | Simulation of up to THREE sets of IEC61850-9-2LE signals shall be possible. | Comply, where Do not Comply provide reasons. |  |  | 3
277 | 4.6.14 | Testing single and three phase transducers | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
278 | a) | Manual testing. | Comply, where Do not Comply provide reasons. |  |  | 3
279 | b) | Automatic testing. | Comply, where Do not Comply provide reasons. |  |  | 3
280 | c) | Generation of sweeps for amplitudes, phase angles, frequency, power (W, VA, Var). | Comply, where Do not Comply provide reasons. |  |  | 3
281 | d) | Individual setting of voltages and currents (ito amplitude, phase angle and frequency) should be possible per test point. | Comply, where Do not Comply provide reasons. |  |  | 3
282 | e) | For single phase transducers phase injected should be selectable. | Comply, where Do not Comply provide reasons. |  |  | 3
283 | f) | L-L voltage transducers to be tested with full three phase voltage system. | Comply, where Do not Comply provide reasons. |  |  | 3
284 | g) | On-line calculation of error (absolute, percentage and full-scale). | Comply, where Do not Comply provide reasons. |  |  | 3
285 | h) | Display of transducer output and absolute, percentage and full-scale errors vs. sweep quantity graphically and/or as table. | Comply, where Do not Comply provide reasons. |  |  | 3
286 | i) | Feedback signal: Low-level analogue voltage (0 - +/- 10V) or current (0 - +/-20mA) | Comply, where Do not Comply provide reasons. |  |  | 3
287 | j) | Automatic assessment of test results (Pass / Fail) with actual results and deviation from nominal. | Comply, where Do not Comply provide reasons. |  |  | 3
288 | 4.6.15 | Power Quality Signal Generator to NRS048-2 / SANS/IEC61000-4-30; SANS/IEC61000-4-7; BS EN/IEC 61000-4-15 and IEC62586 | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
289 | a) | Testing of Power Quality measurement devices as per the NRS 048 standard. | Comply, where Do not Comply provide reasons. |  |  | 3
290 | 4.6.16 | On-line multi-meter measurement function | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
291 | a) | Up to 10 independent inputs | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
292 | b) | Software to display amplitudes and phase angles of ac voltages and currents, symmetrical components of voltages and currents, line to line voltages, frequencies, power (active, reactive and apparent) and cos(phi) independently for each input. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
293 | c) | All ac measurements to be true RMS. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
294 | d) | On-line vector diagram for voltages and currents as well as power. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
295 | e) | On-line measurement of Vdc, Idc and dc power for each input. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
296 | f) | On-line display of harmonics measured both numerically (in a table per harmonic frequency) and graphically in a bar graph. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
297 | 4.6.17 | Transient Recording Function | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
298 | a) | The recording function should enable recordings at the specified sampling frequency. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
299 | b) | Any recordings done should automatically be uploaded to the controller PC. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
300 | c) | The recordings should be saved in COMTRADE format on the PCs hard drive. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
301 | d) | Trigger conditions: amplitude, swell and sag, harmonic, frequency, frequency change, notch as well as any combination of these triggers. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
302 | e) | A pre-trigger buffer should be definable. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
303 | f) | Analysis software should be provided to analyse the recorded wave shapes in terms of analogue wave shapes (RMS and instantaneous values), vector diagram, impedance plots (L-N and L-L fault loops) as well as harmonics. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
304 | g) | For impedance plots, the relays impedance characteristic should be displayable in relation to the impedance trajectory measured by the relay. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
305 | 4.8 | Software Licensing | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
306 | a) | A software license shall cover all operating system versions where the functionality is similar. The future of desktop operating systems is hanging in the balance; end users should have experience with other operating systems such as mobile operating systems. Investment in software should be for the function not the operating system type. A software license shall not depend on a hardware dongle to be used. The License shall be imbedded in the software and the Purchaser’s license shall be embedded in the software package such that any employee of the Purchaser may control a test set of another employee using his / her own computer and test set software, it is accepted modules are registered per test set and that certain functions may not be available to another user. Dedicated software for use with say an IEC61850 package shall also not use a dongle. The license shall be applied such that a user can transfer the license to other computers. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
307 | 4.9 | Local Support | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
308 | a) | Hotline support hours | Comply, where Do not Comply provide reasons. |  |  | 3
309 | b) | Electronic communication support | Comply, where Do not Comply provide reasons. |  |  | 3
310 | c) | Number of units in operation in South Africa | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
311 | d) | No of back-up units available in South Africa | Comply, where Do not Comply provide reasons. |  |  | 3
312 | e) | Calibration facility available in South Africa | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
313 | f) | Guaranteed turnaround time for repairs and calibration | Comply, where Do not Comply provide reasons. |  |  | 3
314 | g) | Product training offered in South Africa | Comply, where Do not Comply provide reasons. |  |  | 3
315 | h) | Warranty on Hardware | Comply, where Do not Comply provide reasons. |  |  | 3
316 | 5 | Test Set configurations | State which test set is offered? (Yes/No) |  | Suppliers reference/Evidence | Weight
317 | 5.1 | Universal three phase test set | Yes/No |  |  | 1
318 | 5.2 | Universal six phase test set | Yes/No |  |  | 3
319 | 5.3 | Universal six phase test set, additional 3 Phase Voltage channel | Yes/No |  |  | 2
320 | 5.4 | Universal Single phase test set | Yes/No |  |  | 1
321 | 6 | Software | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
322 |  | Supplier Complies to all requirements listed under this clause? | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
323 | 6.1 | Basic Software | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
324 | a) | Manual control as defined in item 4 | Comply, where Do not Comply provide reasons. |  |  | 3
325 | b) | Overcurrent and Earth fault module as defined in item 4 | Comply, where Do not Comply provide reasons. |  |  | 3
326 | c) | Frequency and voltage relay testing as defined in item 4 | Comply, where Do not Comply provide reasons. |  |  | 3
327 | Note | (Additional functionality may be listed it)
328 | 6.2 | Intermediate software - includes Basic software items and the additional listed functionality. | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
329 | a) | Ramping module and the pulse ramping module as defined in item 4. | Comply, where Do not Comply provide reasons. |  |  | 3
330 | b) | State Sequencer as defined in item 4 | Comply, where Do not Comply provide reasons. |  |  | 3
331 | c) | Line and Transformer Differential module as defined in item 4 | Comply, where Do not Comply provide reasons. |  |  | 3
332 | d) | Line Distance module as defined in item 4 | Comply, where Do not Comply provide reasons. |  |  | 3
333 | e) | Transducer (single and three phase) testing as defined in item 4 | Comply, where Do not Comply provide reasons. |  |  | 3
334 | f) | Transient playback of Comtrade files as defined in item 4 | Comply, where Do not Comply provide reasons. |  |  | 3
335 | g) | In accordance with IEC61850, GOOSE trigger and simulation as defined in item 4 | Comply, where Do not Comply provide reasons. |  |  | 3
336 | h) | In accordance with IEC61850, Sampled Values simulation as defined in item 4 | Comply, where Do not Comply provide reasons. |  |  | 3
337 | 6.3 | Advanced Advanced modules - This software package shall be separate from the intermediate package, and procured as required by the end user. | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
338 | a) | All software functionality as indicated in the basic and intermediated. | Comply, where Do not Comply provide reasons. |  |  | 3
339 | b) | Dedicated recloser testing modules | Comply, where Do not Comply provide reasons. |  |  | 2
340 | c) | Synchronising modules as defined in item 4 | Comply, where Do not Comply provide reasons. |  |  | 3
341 | d) | Transient recording function as defined in item 4 | Comply, where Do not Comply provide reasons. |  |  | 3
342 | e) | On Line multi-meter function as defined in item 4 | Comply, where Do not Comply provide reasons. |  |  | 2
343 | f) | Power Quality measurement device testing based on NRS048-2 / SANS/IEC61000-4-30; SANS/IEC61000-4-7; BS EN/IEC 61000-4-15 and IEC62586. In essence NRS 048 is the standard followed in South Africa and the Supplier shall state whether the test module is fully compliant with the NRS 048 standard. | Comply, where Do not Comply provide reasons. |  |  | 3
344 | 6.3.1 | Other advanced software - The advanced options as required and where available. | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
345 | a) | Examples of modules/additional hardware not listed are network simulation software, IEC 61850 substation network monitoring/diagnostic tools, MMS communication, simulation-based and distributed protection testing using multiple test sets | State; detail and offer Other Advanced Software Available. |  |  | 3
346 | 7 | Accessories | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
347 | 7.1 | Standard accessory package with a MFT: | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
348 | a) | Each test set to be equipped with a set of 2m; 2 mm² (minimum length) test leads sufficient to populate all inputs/outputs | Comply, where Do not Comply provide reasons. |  |  | 3
349 | b) | A combination test lead with the 4 current and voltage leads, thus 8 leads in total clearly marked for current and voltage application and suitably colour coded. This lead shall be long enough to plug into the current and voltage channels and reach the current and voltage test points on a typical protection or metering scheme. | Comply, where Do not Comply provide reasons. |  |  | 3
350 | c) | Connectors for each lead to allow a sound connection to the devices under test. | Comply, where Do not Comply provide reasons. |  |  | 3
351 | d) | A quality carry bag made of a durable material with carrying strap or other quality case. The bag or container shall easily house all the leads and accessories. | Comply, where Do not Comply provide reasons. |  |  | 3
352 | e) | A hardened transport case with wheels to facilitate easy transport and sufficiently sturdy to be used for multiple shipping events. Lockable. | Comply, where Do not Comply provide reasons. |  |  | 2
353 | f) | The required communication cables | Comply, where Do not Comply provide reasons. |  |  | 3
354 | g) | Manual, the manual may be part of the software module or a software copy. The manual shall be fully detailed. The manual is to cover both the hardware and the software. The manuals may be two separate software documents. The Test set manuals shall preferably be imbedded in the test set control software and have an auto update facility when connected to the Internet. | Comply, where Do not Comply provide reasons. |  |  | 3
355 | h) | Should a software module require a specific accessory to complete or execute a test the test set package shall come equipped with the required accessory and be accordingly priced. | Comply, where Do not Comply provide reasons. |  |  | 3
356 | i) | A calibration certificate. See the details under the calibration section. | Comply, where Do not Comply provide reasons. |  |  | 3
357 | 7.2 | Examples of other Accessories or Services | Schedule A - (Eskom's Specific Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
358 | a) | The listed examples are currently used by the Purchaser, however ALL useful software modules and accessories on offer to be listed and priced in terms of the contract pricing structure and period. | Comply, where Do not Comply provide reasons. |  |  | 3
359 | b) | Clamps, crocodile clips, terminal connectors. | Comply, where Do not Comply provide reasons. |  |  | 3
360 | c) | Clip on current transducers for test sets with current measurement/recording capabilities. | Comply, where Do not Comply provide reasons. |  |  | 3
361 | d) | GPS time synchronisation devices and all accessories required to execute the time synchronisation function. All offerings such as dedicated units and PTP based units. | Comply, where Do not Comply provide reasons. |  |  | 3
362 | e) | ARC sensor light emitter. (Electrical arc protection scheme testing) | Comply, where Do not Comply provide reasons. |  |  | 2
363 | f) | Sensors to accommodate measurement of all types of meters. (Pulsing LED signals) | Comply, where Do not Comply provide reasons. |  |  | 3
364 | g) | Additional communications cables and other specialised cables. | Comply, where Do not Comply provide reasons. |  |  | 3
365 | h) | Software modules | Comply, where Do not Comply provide reasons. |  |  | 3
366 | i) | Calibration service to SANAS requirement, priced for all test set offerings including return shipping. | Comply, where Do not Comply provide reasons. |  |  | 3
367 | j) | Repair / Calibration return shipping, however this is usually included in the repair / calibration cost and shall be the preferred method. | Comply, where Do not Comply provide reasons. |  |  | 3
368 | k) | High voltage testing protective devices (e.g. Power Transformer testing) such as metrosil / varistor type protection | Comply, where Do not Comply provide reasons. |  |  | 3
369 | l) | Switch boxes | Comply, where Do not Comply provide reasons. |  |  | 3
370 | m) | Test set controllers with on board software (Not a PC) | Comply, where Do not Comply provide reasons. |  |  | 3
371 | n) | Dedicated cables to match reclosers etc. | Comply, where Do not Comply provide reasons. |  |  | 1
372 | o) | Dedicated recloser test software | Comply, where Do not Comply provide reasons. |  |  | 2
373 | p) | Hardware upgrades | Comply, where Do not Comply provide reasons. |  |  | 3
374 | q) | Any other useful accessory or software module | Comply, where Do not Comply provide reasons. |  |  | 3
Company Name: |  |  |  | Date
Representative Name |  |  |  | Signature

## A2-A&B Technical Support Requir
A&B SCHEDULE FOR MULTI-FUNCTION CONTROL PLANT TEST SET SPECIFICATION 240-170000773 PROTECTION MFT SUPPORT REQUIREMENTS
Schedule A: Purchaser specific requirements
Schedule B: Guarantees and technical particulars of equipment offered
When completing the Schedule of Requirements A/B Suppliers are requested to take cognisance of the following: • It is expected of Suppliers to state clearly, for each clause that requires a statement of compliance in the Schedule of Requirements A/B, either “Comply” or “Do not comply”. • If a clause in the Schedule of Requirements A/B requires a statement of compliance and additional information e.g. a description, reasons etc., Suppliers are requested to state clearly either “Comply – followed by any information deemed appropriate by the Supplier” or “Do not comply – followed by any information deemed appropriate by the Supplier”. • If a clause in the Schedule of Requirements A/B requires information only, e.g. describe application etc., Suppliers are at liberty to provide any information deemed appropriate. Supporting information, such as specification sheets, instruction manuals, test reports and certificates should be referenced where applicable. Electronic copies of the information referenced shall be provided and can be hyperlinked where possible to assist with the evaluation. The Supplier shall provide detailed evidence for each answer specifically stating the exact location of the evidence in the supporting documentation i.e. document name and reference, page number, paragraph number It is critical that references to supporting documentation (document title, page number and clause / paragraph no.) that confirm compliance shall be clearly indicated on the Technical Schedules in the "Reference" column.
Clause of 240-170000773 | Description | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
8.1 | Calibration
a) | Ability of Supplier to confirm fully to Clause 8.1 | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
8.2 | Repair | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Ability of Supplier to confirm fully to Clause 8.2 | Comply, where Do not Comply provide reasons. |  |  | 3
8.3 | Annual Report | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Ability of Supplier to confirm fully to Clause 8.3 | Comply provide evidence, where Do not Comply provide reasons. |  |  | 2
9 | Training | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Ability of Supplier to confirm fully to Clause 9 | Comply, where Do not Comply provide reasons. |  |  | 3
9.1 | Basic Training | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Is the relevant training material available and ability to conform to requirements? | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
9.2 | Intermediate Training | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Is the relevant training material available and ability to conform to requirements? | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
9.3 | Metering and Measurements Training | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Is the relevant training material available and ability to conform to requirements? | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
9.4 | Advanced Training | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Is the relevant training material available and ability to conform to requirements? | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
9.5 | Specialist Training | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Is the relevant training material available and ability to conform to requirements? | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
9.6 | Trainee Evaluation | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Ability of supplier to conform | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
10 | Support | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Local repair and calibration with a turnaround time typically better than 14 days. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
b) | Formalised training courses as described. Training shall take place at or near the Purchaser‘s offices around the country, when sufficient student numbers justify this; or virtual training as required. | Comply, where Do not Comply provide reasons. |  |  | 3
c) | Technical support shall be available via email and telephonically or any other form of electronic communication by technical competent support staff who is within 2 hours of South African time zone. Technical competent would mean personnel with tertiary Electrical engineering qualifications and having hands on experience testing protection IED’s that are used in the Purchaser’s division (Transmission.) | Comply, where Do not Comply provide reasons. |  |  | 3
e) | Lastly the Supplier shall provide proof of the required support functions. This will be a list of technical staff and a short CV of each member, courses provided over the past years and to which companies. Physical examples of automated templates and the end customer. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
11 | Automated Test Templates | Schedule A - (Eskom's minimum technical Requirement) | Schedule B(Supplier's statement of compliance) | Suppliers reference/Evidence | Weight | Technical Score
a) | Capability and capacity to develop automated testing templates on request. | Comply provide evidence, where Do not Comply provide reasons. |  |  | 3
b) | Compatibility with the Purchaser’s existing automated test templates is required. Alternatively, Suppliers can submit proposals to migrate existing automated test templates to a format that is compatible with their proposed products. | Comply or submit proposals |  |  | 3
|  |  | Total Risk and Support Score |  | 47
Company Name: |  |  |  | Date
Representative Name |  |  |  | Signature

## B1 - Practical Evaluation
PRACTICAL DEMONSTRATION FOR PROTECTION MULTI-FUNCTION CONTROL PLANT TEST SET SPECIFICATION 240-170000773
Number | Demonstration | Weighting | Supplier Score | Comments
1 | Protection Testing of electromechanical and electronic schemes | 10
2 | Protection Testing of Microprocessor based schemes | 10
3 | Protection Testing of Microprocessor based schemes utilizing IEC61850 and associated protocols. | 10
4 | Testing of measurement/metering devices | 10
5 | Ability of test set to provide test reports usable to Eskom, eg in pdf | 10
6 | Capability to import a Eskom setting file in the test template. | 10
7 | Demonstrate fault play back, with a file provided by the Eskom team. | 5
8 | Demonstrate the use of IEC61850 network commissioning tools. | 10
9 | Demonstration of requirements based on a single or multiple test sets to achieve the required functionality. | 5
10 | Demonstration of synchronised testing among multiple test sets with the same manufacturer. | 6
11 | Demonstration of synchronised testing among multiple test sets with different manufacturer. | 4
12 | Demonstration of an automated test template utilising IEC 61850 protocol. | 10
Company Name: |  |  | Date
Eskom Evaluator Names |  |  | Eskom Evaluator Signature

## Deviation Schedule
Deviation Schedule
Any deviations from the 240-170000773 specification shall be listed below with reasons for deviation. In addition, evidence shall be provided that the proposed deviation will at least be more cost effective than that specified by Eskom.
Item | Clause | Proposed Deviation
Company Name
Representative Name
Date
Signature